## Software Specification
requirement ID | Area | Text | Priority(Low, Med. High) | Allocation | Rationale | Clinical system | Bench top system | Software/ Hardware
REQ1 | Data - Header file | Header file with system parameters: list | High | XZ | Make sure that everything is available for future use | X | X | Software
REQ2 | Data - IRF Alignment save | Make sure that the full waveform of the IRF alignment step is saved | High | XZ | used for intensity ratio calibration during offline data processing | X |  | Software
REQ3 | Data - TDMS file format | Define output data file format | High | XZ | Will be different since we have individual ADC channels for each instrument channel | X | X | Software
REQ4 | Init - Display reference WF | Display of reference standard waveform, absolute and normalized | High | XZ | Verify amount of signal and relative channel intensity | X | X | Software
REQ5 | Init - Load last use parameters | Default loading of last used parameters | High | XZ | Minimize user input, quick resume in case of software crash | X | X | Software
REQ6 | Init - BG measurement steps | include step explicitely in background measurement, define FoM for sampling port adjustment | Medium | XZ | During background measurement, minimize proximal background, maximize distal end. Maximize FoM reduces subjectivity | X |  | Software
7 | Init - Probe switch status | Ability to probe switch all status: requires 2 circuits, actuation/ status | Medium | XZ | Provide user guidance about next step, state change following switch actuation by user | X |  | Software/ Hardware
REQ8 | Init - Interlock T7 | ability to disable interlock/start laser with T7? | Low | XZ | both T7 and pushbutton can achieve the same funcitonality. which option best suit safty compliance? | X |  | Software/ Hardware
REQ9 | Main - Critical parameters | Specific “maintenance” page where critical parameters are accessed | High | XZ | No erroneous modification of key parameters | X |  | Software
REQ10 | Main - APD response curve | APD response curve calibration | High | XZ | May be easier to include in the main software than having to use an external code to measure it and then import the curves | X | X | Software
REQ11 | Main - Reference WF | Store new reference waveform | High | XZ | Important that it is not done by mistake or we may lose track of the reference scan. | X | ? | Software
REQ12 | Main - Gain control parameters | gain control parameters | High | XZ | May be improved in the future | X | X | Software
REQ13 | Main - Laser reprate adjustment | increase laser rep rate for accurate power measurement | High | XZ | Low reliability of power meter at low rep rate. Needs to be in maintenance mode to not have issue with certification? | X | ? | Software
REQ14 | Main - Hardware IO status | Visualization of hardware IO states: toggle shutter, maintenance verify pedal operation, etc... | Low | XZ |  | X | X | Software
REQ15 | Main - Load all parameter | Ability to load all parameters from setting file | High | XZ | Setting file provided by us, possibility to easily apply new parameter sets if needed | X |  | Software
REQ16 | Main - Turn on/off laser | Ability to easily turn on/off each light source | High | XZ | used for alignment and intensity calibration | X |  | Software
REQ17 | Use - Data display | Display waveforms, chosen overlay and SNR overlay | High | XZ | Supports acquisition of high quality data | X | ? | Software
REQ18 | Use - Chime | Chime when acquisition is proceeding | Low | XZ |  |  |  | Software
REQ19 | Use - Clear system status indication | Clear indication of system status | High | XZ | inform user whether the system is ready for data acquisition, required for compliance | X | X | Software
20 | Use - Data visualization | Data visualization: available for newly acquired run as well as previous runs | Low | XZ | Verify data quality | X | X | Software
21 | Use - Operation mode selection | Easy selection of operation mode | High | XZ |  | X | X | Software
REQ22 | Use - Scan option | Suspend/ resume/ abort scan options: stop laser, stop motion | High | XZ | Flexible and safe use, common for machine tools: if the emergency stop shuts everything off, less likely to use it because of the lost time | X | X | Software
REQ23 | Use - Scout Scan | Scout scan option | Medium | XZ | Easier workflow than modify the scanning parameters to coarser values |  | X | Software
REQ24 | Use - Color coding button | Color coding of items modified during this session | Medium | XZ | Make clear when automatically loaded parameters are used | X | X | Software
REQ25 | Use - Show both gain and control voltage | Show system both system gain and control voltage | Medium | XZ | Show both control voltage and APD gain for user | X | X | Software
REQ26 | Use - System health index | Parameter to show system health status based on refrence | High | XZ | Parameter to show system health status based on refrence | X | X | Software
REQ27 | Main - Aiming Beam laser 0Hz operation | Ability to run aiming beam laser at 0Hz for alignment | High | XZ | Be able to run aiming beam laser at 0Hz for system alignment | x | x | Software
REQ28 | Main - check PWM in valid range | PWM can not go below 5Hz | High | XZ | throw error when PWM setting is not valid, i.e 1 - 5Hz | x | x | Software
REQ31 | Main - IRF measurement | Procedure and user interface to guide acquisition of single detector IRF, and save in reference folder | High | XZ | define a list of gain/bias setpoints (not necessarily evenly spaced, may need more at low gain), specify user to adjust the laser attenuation to obtain a given amplitude for each point and channel and acquire IRF from laser, cue to save. Suggestion: possibility to do all detectors in any order to optimize laser adjustment | x | x | Software
REQ32 | Use - Blinking Green LED during acquisition | Blinking Green LED during quistion | Medium | XZ | Blinking Green LED during quistion | x | x | Software
REQ33 | Use - Aimging Beam momentery ON | Aimging Beam momentery ON | High | XZ | Turn aimging beam ON when front panel button is pressed | x | x | Software
REQ34 | Use - 355nm Laser momentery ON | Laser momentery ON | High | XZ | Turn 355nm laser on when front panel button is pressed
REQ35 | Use - 5min acquisition timer | 5min acquisition timer | High | XZ | Maximum acquisition time is 5min, stop if longer than 5min | x | x | Software
REQ36 | Use - Show ready screen after LV is ready | Show ready screen on AB computer after system finishes initialization | High | XZ | Show ready screen after LV inititialization | x | x | Software
REQ37 | Main - unplugging the pedal does not trigger an acquisition | By design the circuit opens when the jack is inserted, which triggers an acquisition, this is a safety concern | Medium | XZ | Possible solution: timer on "open circuit", to make sure that a short open circuit (as when pulling the jack) does not trigger an acquisition. Maybe aim for 0.5s and check how this feels? (effectively address the issue, without making the system too slow to respond) | x

## Hardware Specification
ID | Area | Text | Priority(Low, Med. High) | Allocation | Rationale | Clinical system | Bench top system | Software/ Hardware
27 | Init - Sampling port adjustment | Sampling port adjustment | High | XZ | During background measurement, minimize proximal background, maximize distal end
28 | Main - UV attenuation | UV Laser attenuation adjustment | High | XZ | Combination of ND filter and blade position

## Safety Specification
requirement ID | Area | Text | Priority(Low, Med. High) | Allocation | Rationale | Clinical system | Bench top system | Software/ Hardware | Count(for sorting)
REQ29 | Safty - Laser rep rate failure | Check laser rep rate, if not with in 1%, stop system | High | XZ | Stop system if single point failure of the laser is detected | X | X |  | 29
REQ30 | Safty - Laser interlock failure | Safely shutdown system when interlock opens | High | XZ | If interlock is open, stop acquistion and safely shut down system | X | X |  | 30

## Key system parameters
Parameter name | Parameter description | Safety related | Related to spec #
| Aiming beam current | Y
| Aiming beam modulation frequency | N
| Laser repetition rate | Y
| number of averaging | N
| Gain control target voltage | N
| Gain control proportional gain | N
| Gain control derivative gain | N
| "proximal" time window for sampling port alignment FoM | N | 22
| "Distal" time window for sampling port alignment FoM | N | 22