## Software Specification
requirement ID | Area | Text | Priority(Low, Med. High) | Allocation | Rationale | Clinical system | Bench top system | Software/ Hardware
REQ1 | Data - Header file | Header file with system parameters: list | High | XZ | Make sure that everything is available for future use | X | X | Software
REQ2 | Data - IRF Alignment save | Make sure that the full waveform of the IRF alignment step is saved | High | XZ | used for intensity ratio calibration during offline data processing | X |  | Software
REQ3 | Data - TDMS file format | Define output data file format | High | XZ | Will be different since we have individual ADC channels for each instrument channel | X | X | Software
REQ4 | Init - Display reference WF | Display of reference standard waveform, absolute and normalized | High | XZ | Verify amount of signal and relative channel intensity | X | X | Software
REQ5 | Init - Load last use parameters | Default loading of last used parameters | High | XZ | Minimize user input, quick resume in case of software crash | X | X | Software
REQ6 | Init - BG measurement steps | include step explicitely in background measurement, define FoM for sampling port adjustment | Medium | XZ | During background measurement, minimize proximal background, maximize distal end. Maximize FoM reduces subjectivity | X |  | Software
7 | Init - Probe switch status | Ability to probe switch all status: requires 2 circuits, actuation/ status | Medium | XZ | Provide user guidance about next step, state change following switch actuation by user | X |  | Software/ Hardware
REQ8 | Init - Interlock T7 | ability to disable interlock/start laser with T7? | Low | XZ | both T7 and pushbutton can achieve the same funcitonality. which option best suit safty compliance? | X |  | Software/ Hardware
REQ9 | Main - Critical parameters | Specific “maintenance” page where critical parameters are accessed | High | XZ | No erroneous modification of key parameters | X |  | Software
REQ10 | Main - APD response curve | APD response curve calibration | High | XZ | May be easier to include in the main software than having to use an external code to measure it and then import the curves | X | X | Software
REQ11 | Main - Reference WF | Store new reference waveform | High | XZ | Important that it is not done by mistake or we may lose track of the reference scan. | X | ? | Software
REQ12 | Main - Gain control parameters | gain control parameters | High | XZ | May be improved in the future | X | X | Software
REQ13 | Main - Laser reprate adjustment | increase laser rep rate for accurate power measurement | High | XZ | Low reliability of power meter at low rep rate. Needs to be in maintenance mode to not have issue with certification? | X | ? | Software
REQ14 | Main - Hardware IO status | Visualization of hardware IO states: toggle shutter, maintenance verify pedal operation, etc... | Low | XZ |  | X | X | Software
REQ15 | Main - Load all parameter | Ability to load all parameters from setting file | High | XZ | Setting file provided by us, possibility to easily apply new parameter sets if needed | X |  | Software
REQ16 | Main - Turn on/off laser | Ability to easily turn on/off each light source | High | XZ | used for alignment and intensity calibration | X |  | Software
REQ17 | Use - Data display | Display waveforms, chosen overlay and SNR overlay | High | XZ | Supports acquisition of high quality data | X | ? | Software
REQ18 | Use - Chime | Chime when acquisition is proceeding | Low | XZ |  |  |  | Software
REQ19 | Use - Clear system status indication | Clear indication of system status | High | XZ | inform user whether the system is ready for data acquisition, required for compliance | X | X | Software
20 | Use - Data visualization | Data visualization: available for newly acquired run as well as previous runs | Low | XZ | Verify data quality | X | X | Software
21 | Use - Operation mode selection | Easy selection of operation mode | High | XZ |  | X | X | Software
REQ22 | Use - Scan option | Suspend/ resume/ abort scan options: stop laser, stop motion | High | XZ | Flexible and safe use, common for machine tools: if the emergency stop shuts everything off, less likely to use it because of the lost time | X | X | Software
REQ23 | Use - Scout Scan | Scout scan option | Medium | XZ | Easier workflow than modify the scanning parameters to coarser values |  | X | Software
REQ24 | Use - Color coding button | Color coding of items modified during this session | Medium | XZ | Make clear when automatically loaded parameters are used | X | X | Software
REQ25 | Use - Show both gain and control voltage | Show system both system gain and control voltage | Medium | XZ | Show both control voltage and APD gain for user | X | X | Software
REQ26 | Use - System health index | Parameter to show system health status based on refrence | High | XZ | Parameter to show system health status based on refrence | X | X | Software
REQ27 | Main - Aiming Beam laser 0Hz operation | Ability to run aiming beam laser at 0Hz for alignment | High | XZ | Be able to run aiming beam laser at 0Hz for system alignment | x | x | Software
REQ28 | Main - check PWM in valid range | PWM can not go below 5Hz | High | XZ | throw error when PWM setting is not valid, i.e 1 - 5Hz | x | x | Software
REQ31 | Main - IRF measurement | Procedure and user interface to guide acquisition of single detector IRF, and save in reference folder | S | XZ | define a list of gain/bias setpoints (not necessarily evenly spaced, may need more at low gain), specify user to adjust the laser attenuation to obtain a given amplitude for each point and channel and acquire IRF from laser, cue to save. Suggestion: possibility to do all detectors in any order to optimize laser adjustment | x | x | Software
REQ32 | Use - Blinking Green LED during acquisition | Blinking Green LED during quistion | Medium | XZ | Blinking Green LED during quistion | x | x | Software
REQ33 | Use - Aimging Beam momentery ON | Aimging Beam momentery ON | High | XZ | Turn aimging beam ON when front panel button is pressed | x | x | Software
REQ34 | Use - 355nm Laser momentery ON | Laser momentery ON | High | XZ | Turn 355nm laser on when front panel button is pressed
REQ35 | Use - 5min acquisition timer | 5min acquisition timer | High | XZ | Maximum acquisition time is 5min, stop if longer than 5min | x | x | Software
REQ36 | Use - Show ready screen after LV is ready | Show ready screen on AB computer after system finishes initialization | High | XZ | Show ready screen after LV inititialization | x | x | Software

## Hardware Specification
ID | Area | Text | Priority(Low, Med. High) | Allocation | Rationale | Clinical system | Bench top system | Software/ Hardware
27 | Init - Sampling port adjustment | Sampling port adjustment | High | XZ | During background measurement, minimize proximal background, maximize distal end
28 | Main - UV attenuation | UV Laser attenuation adjustment | High | XZ | Combination of ND filter and blade position

## Safety Specification
requirement ID | Area | Text | Priority(Low, Med. High) | Allocation | Rationale | Clinical system | Bench top system | Software/ Hardware | Count(for sorting)
REQ29 | Safty - Laser rep rate failure | Check laser rep rate, if not with in 1%, stop system | High | XZ | Stop system if single point failure of the laser is detected | X | X |  | 29
REQ30 | Safty - Laser interlock failure | Safely shutdown system when interlock opens | High | XZ | If interlock is open, stop acquistion and safely shut down system | X | X |  | 30

## Key system parameters
Parameter name | Parameter description | Safety related | Related to spec #
| Aiming beam current | Y
| Aiming beam modulation frequency | N
| Laser repetition rate | Y
| number of averaging | N
| Gain control target voltage | N
| Gain control proportional gain | N
| Gain control derivative gain | N
| "proximal" time window for sampling port alignment FoM | N | 22
| "Distal" time window for sampling port alignment FoM | N | 22